RADIATION SAFETY STANDARD PROTOCOL REQUIREMENTS: IODINATIONS

Pre-Iodination

1.	Prior to handling, or observing work with a volatile form of radioiodine for the first time, each individual will have a baseline thyroid bioassay.  Thereafter, protocol users will respond promptly to Division of Radiation Safety (DRS) requests for thyroid bioassays.

- Iodinations will be performed in an approved iodination box located in a chemical fume hood that has been certified within the last 12 months and has a green calibration sticker indicating that the hood is suitable for work with toxic compounds or low-level radioactive material.

- The charcoal filter located on the top of the iodination box will be replaced as needed per HP evaluation but no longer than every 4 years. The evaluation will be based on factors that include, but not limited to: increased radioactivity in previous effluent samples; exposure from the filter is not ALARA; the manufacturer recommends changing the filter under specific conditions (e.g. a filter used with an exhaust system under continuous operation.)

- A sampling station to monitor exhaust air effluent from the fume hood containing the iodination box must be installed by DRS prior to the performance of any protocol work.

- Protocol users shall notify the DRS at least 1 business day in advance of each iodination by e-mailing Iodinations@mail.nih.gov .  Researchers shall provide the date, approximate start and end times, and the room number where the iodination will be performed.

- Prior to beginning of each iodination, the individual user will ensure that the exhaust fan in the iodination box above the charcoal filter is working by testing the airflow through the arm ports.

During Iodination

- Dosimetry is required if the activity handled exceeds the following:

| Isotope | Whole-body mCi-min/yr | Extremity mCi-min/yr |
| --- | --- | --- |
| I-123 | 35,700 | 1,750 |
| I-124 | 4,300 | 45 |
| I-125 | Not required | Not required |
| I-131 | 11,520 | 540 |

- Work with volatile radioactive material will stop, the hood sash or door closed, and DRS notified immediately if a malfunction of the iodination hood is suspected or the fume hood has a red sticker on it.

- The protocol user shall collect a breathing zone air sample throughout the iodination.  Air flow through the sampler will be adjusted to 4 liters per minute (LPM).

- Tongs, forceps and other remote handling devices will be used when opening or otherwise manipulating source containers.

- Waste containers, source vials, and other concentrated sources may require lead shielding. The goal is to reduce the dose rate to less than 5 mR/hr @ 30 cm if possible:

| Isotope | mR/hr-mCi @ 30 cm point source | mR/hr-mCi contact 5 ml plastic syringe | Skin dose rate R/hr-µCi/cm2 | HVL - PB mm / inches |
| --- | --- | --- | --- | --- |
| I-123 | 1.91E+0 | 2.24E+3 | 4.69E-2 | 1 / < 1/16 |
| I-124 | 7.07E+0 | 3.96E+4 | 1.66E+0 | 8 / 5/16 |
| I-125 | 1.44E+0 | 2.29E+3 | 4.75E-2 | <1 / <1/16 |
| I-131 | 2.70E+0 | 4.18E+3 | 5.29E+0 | 3 / 1/8 |

- Stock solutions and concentrated (high-activity, low-volume) waste solutions will be segregated from other radioactive waste and stored in tightly-capped containers.  Opening of such containers will only be performed inside of the iodination box with the exhaust fan for the box in the "on" condition.

- Other waste containing I-123, I-124, I-125 and/or I-131 will be disposed with other radionuclides having half-lives of less than 100 days.

Post-Iodination

- Following an iodination, hands, arms, clothing, and shoes will be monitored for contamination using a low energy sodium iodide scintillator for I-125 or a Geiger counter for I-123, I-124 or I-131.  The work area, including the floor in front of the iodination hood will be smeared for contamination.  If any smears are above 2200 dpm per 100 cm2 the affected areas will be immediately decontaminated; additional smears will be taken to ensure that contamination levels are less than 2200 dpm per 100 cm2 in the work area and less than 220 dpm per 100 cm2 in unrestricted areas.

- A log book will be maintained with the following information: date, name of user, DRS number, radionuclide, activity in source vial, activity used, verification of a green sticker affixed to the hood, verification that the hood has been calibrated within 1 year, results from the post-iodination meter and smear surveys, verification that form NIH 88-17 was filled out, and an indication that DRS was contacted for the breathing zone sample.

- When the iodination procedure is complete, the protocol user shall remove the breathing zone sampler, and e-mail Iodinations@mail.nih.gov to arrange for collection of the breathing zone sampler and completed air sampling report form (NIH 88-17) including: all iodinators; the radionuclide; the activity in the source vial; and the activity used.

- Any medical-pathological waste that is contaminated with radioiodine will be disposed of through the radioactive waste service after any infectious agents have been deactivated using a non-chlorine-based disinfectant such as Wescodyne.

- Radioactive waste pickups will be scheduled promptly following iodinations.

Other

- Any spill or personnel contamination resulting from protocol work will be reported to the DRS as soon as possible.

- Any positive breathing zone sample, or individual involved in a spill or personnel contamination incident associated with the protocol must have a thyroid bioassay between 8 and 72 hours following the incident.  Call 301-496-5774 to schedule the bioassay.

- Any airborne effluent sample that exceeds 1% of the constraint to members of the public needs to be investigated. No further iodination will be performed until the iodination charcoal filter is changed.

- All radioactive materials in use or storage, including waste, must be secured from unauthorized removal or access when unattended.

- The DRS must be notified before any modifications are made to the iodination hood, protocol laboratory, or protocol procedures.

- Women who choose to declare their pregnancies are advised not to perform iodinations.